## Introduction
1 | This study was performed at several US ARCHITECT sites utilizing the Reagent Carryover Corrective Action Procedures located in Section 10 of the ARCHITECT Systems Operations Manual or the Clinical Chemistry Applications Guide.
2 | Most runs were performed August 17 through August 20, 2015 by Abbott Technical Assay Specialists personnel.
3 | Not all assays are available and they are not in alphabetical order.
4 | Representative data, individual laboratory results may vary.
5 | Additional RCC studies are being conducted and will be disseminated when available.
6 | Non-Abbott applications are marked with an asterisk*.
7 | Our recommendation is to retain suggested Suggested SmartWashes from attachment Acetaminpehn 3R11 Recommended SmartWashes and located at the Sekisui customer website (http://www.sekisuidiagnostics.com/landing/acetaminophen).
8 | Also, if the extent of carryover for any of the assays in this representative data set is outside the acceptable limits for your laboratory, Architect Operations Manual Section 3 provides information on setting up additional Smartwashes.
9 | Replicate 1 = Value for cup 5, rep 1 / Target = Mean of cup 2, all replicates / See Calculation Below:

## List of Assays Not Tested
Assay's Missing From Study
1 | Acid Phospahtase
2 | AGP
3 | A-AST
4 | A-ALT
5 | Apo A
6 | Apo B
7 | *Bile Acid
8 | Cannabinoid (THC)
9 | *CHE (Cholinesterase)
10 | *CKMB
11 | *Copper
12 | *Cystatin C
13 | *D-Dimer
14 | *Dibucaine CHE
15 | *Digitoxin
16 | Ecstasy
17 | Ferritin
18 | *Fructosamine
19 | HbA1c (LN 4P52 and 2K96)
20 | HIL
21 | *HBDH
22 | *Kappa
23 | *Lambda
24 | Lpa
25 | *Mag (LN 3P68)
26 | *Myoglobin
27 | *Pancreatic Amylase
28 | PCPQ (acetaminophen into Acet)
29 | Quinidine
30 | *RF (LN 6K44)
31 | UIBC
*Assay's Available Outside U.S.

## ACET_Acet_AlbG_AlkP_ALT
| Carrier | 1 |  |  | Carrier | 2 |  |  | Carrier | 3
Donor |  | Recipient |  | Donor |  | Recipient |  | Donor |  | Recipient
(Cup 4) | ACET | (Cup 5) | ACET | (Cup 4) | ACET | (Cup 5) | ACET | (Cup 4) | ACET | (Cup 5) | Acet
List Number | 3R11 |  | 3R11 | List Number | 3R11 |  | 3R11 | List number | 3R11 |  | 2K99
Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results
Rep 1 | 146 | Rep 1 | 146 | Rep 1 | 145 | Rep 1 | 145 | Rep 1 | 131 | Rep 1 | 131
Rep 2 | 145 | Rep 2 | 145 | Rep 2 | 145 | Rep 2 | 145 | Rep 2 | 131 | Rep 2 | 132
Rep 3 | 145 | Rep 3 | 145 | Rep 3 | 145 | Rep 3 | 145 | Rep 3 | 131 | Rep 3 | 132
Rep 4 | 146 | Rep 4 | 145 | Rep 4 | 145 | Rep 4 | 145 | Rep 4 | 131 | Rep 4 | 132
Rep 5 | 146 | Rep 5 | 145 | Rep 5 | 144 | Rep 5 | 145 | Rep 5 | 131 | Rep 5 | 131
Target | 145.6 | % Diff. | 0.27473 | Target | 144.8 | % Diff. | 0.13812 | Target | 131 | % Diff. | 0
(mean) |  | (Rep 1) |  | (mean) |  | (Rep 1) |  | (mean) |  | (Rep 1)
| Carrier | 4 |  |  | Carrier | 5 |  |  | Carrier | 6
Donor |  | Recipient |  | Donor |  | Recipient |  | Donor |  | Recipient
(Cup 4) | Acet | (Cup 5) | ACET | (Cup 4) | ACET | (Cup 5) | AlbG | (Cup 4) | AlbG | (Cup 5) | ACET
List Number | 2K99 |  | 3R11 | List Number | 3R11 |  | 7D53 | List Number | 7D53 |  | 3R11
Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results
Rep 1 | 144 | Rep 1 | 145 | Rep 1 | 4.1 | Rep 1 | 4.1 | Rep 1 | 145 | Rep 1 | 145
Rep 2 | 144 | Rep 2 | 145 | Rep 2 | 4.1 | Rep 2 | 4.1 | Rep 2 | 145 | Rep 2 | 145
Rep 3 | 144 | Rep 3 | 145 | Rep 3 | 4.1 | Rep 3 | 4 | Rep 3 | 145 | Rep 3 | 144
Rep 4 | 144 | Rep 4 | 145 | Rep 4 | 4.1 | Rep 4 | 4.1 | Rep 4 | 144 | Rep 4 | 144
Rep 5 | 146 | Rep 5 | 145 | Rep 5 | 4.1 | Rep 5 | 4.1 | Rep 5 | 144 | Rep 5 | 143
Target | 144.4 | % Diff. | 0.41551 | Target | 4.1 | % Diff. | 0 | Target | 144.6 | % Diff. | 0.27663
(mean) |  | (Rep 1) |  | (mean) |  | (Rep 1) |  | (mean) |  | (Rep 1)
| Carrier | 7 |  |  | Carrier | 8 |  |  | Carrier | 9
Donor |  | Recipient |  | Donor |  | Recipient |  | Donor |  | Recipient
(Cup 4) | ACET | (Cup 5) | AlkP | (Cup 4) | AlkP | (Cup 5) | ACET | (Cup 4) | ACET | (Cup 5) | ALT
List number | 3R11 |  | 7D55 | List Number | 7D55 |  | 3R11 | List Number | 3R11 |  | 7D56
Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results
Rep 1 | 310 | Rep 1 | 310 | Rep 1 | 143 | Rep 1 | 144 | Rep 1 | 195 | Rep 1 | 194
Rep 2 | 310 | Rep 2 | 309 | Rep 2 | 143 | Rep 2 | 143 | Rep 2 | 194 | Rep 2 | 194
Rep 3 | 309 | Rep 3 | 307 | Rep 3 | 143 | Rep 3 | 144 | Rep 3 | 194 | Rep 3 | 193
Rep 4 | 310 | Rep 4 | 308 | Rep 4 | 143 | Rep 4 | 143 | Rep 4 | 194 | Rep 4 | 193
Rep 5 | 311 | Rep 5 | 307 | Rep 5 | 142 | Rep 5 | 142 | Rep 5 | 194 | Rep 5 | 193
Target | 310 | % Diff. | 0 | Target | 142.8 | % Diff. | 0.84034 | Target | 194.2 | % Diff. | -0.10299
(mean) |  | (Rep 1) |  | (mean) |  | (Rep 1) |  | (mean) |  | (Rep 1)

## ALT_Amy_AST_BilD_BilT
| Carrier | 10 |  |  | Carrier | 11 |  |  | Carrier | 12
Donor |  | Recipient |  | Donor |  | Recipient |  | Donor |  | Recipient
(Cup 4) | ALT | (Cup 5) | ACET | (Cup 4) | ACET | (Cup 5) | Amy | (Cup 4) | Amy | (Cup 5) | ACET
Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results
Rep 1 | 143 | Rep 1 | 144 | Rep 1 | 312 | Rep 1 | 314 | Rep 1 | 144 | Rep 1 | 144
Rep 2 | 144 | Rep 2 | 144 | Rep 2 | 310 | Rep 2 | 313 | Rep 2 | 143 | Rep 2 | 143
Rep 3 | 143 | Rep 3 | 144 | Rep 3 | 309 | Rep 3 | 315 | Rep 3 | 143 | Rep 3 | 144
Rep 4 | 143 | Rep 4 | 144 | Rep 4 | 312 | Rep 4 | 313 | Rep 4 | 144 | Rep 4 | 142
Rep 5 | 142 | Rep 5 | 144 | Rep 5 | 313 | Rep 5 | 314 | Rep 5 | 143 | Rep 5 | 143
Target | 143 | % Diff. | 0.6993 | Target | 311.2 | % Diff. | 0.89974 | Target | 143.4 | % Diff. | 0.41841
(mean) |  | (Rep 1) |  | (mean) |  | (Rep 1) |  | (mean) |  | (Rep 1)
| Carrier | 13 |  |  | Carrier | 14 |  |  | Carrier | 15
Donor |  | Recipient |  | Donor |  | Recipient |  | Donor |  | Recipient
(Cup 4) | ACET | (Cup 5) | AST | (Cup 4) | AST | (Cup 5) | ACET | (Cup 4) | ACET | (Cup 5) | BilD
Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results
Rep 1 | 255 | Rep 1 | 254 | Rep 1 | 145 | Rep 1 | 145 | Rep 1 | 2.8 | Rep 1 | 2.8
Rep 2 | 255 | Rep 2 | 254 | Rep 2 | 145 | Rep 2 | 145 | Rep 2 | 2.8 | Rep 2 | 2.8
Rep 3 | 255 | Rep 3 | 253 | Rep 3 | 144 | Rep 3 | 145 | Rep 3 | 2.8 | Rep 3 | 2.8
Rep 4 | 257 | Rep 4 | 253 | Rep 4 | 144 | Rep 4 | 145 | Rep 4 | 2.8 | Rep 4 | 2.8
Rep 5 | 255 | Rep 5 | 250 | Rep 5 | 144 | Rep 5 | 145 | Rep 5 | 2.7 | Rep 5 | 2.7
Target | 255.4 | % Diff. | -0.54816 | Target | 144.4 | % Diff. | 0.41551 | Target | 2.78 | % Diff. | 0.71942
(mean) |  | (Rep 1) |  | (mean) |  | (Rep 1) |  | (mean) |  | (Rep 1)
| Carrier | 16 |  |  | Carrier | 17 |  |  | Carrier | 18
Donor |  | Recipient |  | Donor |  | Recipient |  | Donor |  | Recipient
(Cup 4) | BilD | (Cup 5) | ACET | (Cup 4) | ACET | (Cup 5) | BilT | (Cup 4) | BilT | (Cup 5) | ACET
Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results
Rep 1 | 145 | Rep 1 | 145 | Rep 1 | 7.3 | Rep 1 | 7.3 | Rep 1 | 146 | Rep 1 | 146
Rep 2 | 145 | Rep 2 | 144 | Rep 2 | 7.2 | Rep 2 | 7.3 | Rep 2 | 145 | Rep 2 | 147
Rep 3 | 145 | Rep 3 | 145 | Rep 3 | 7.2 | Rep 3 | 7.2 | Rep 3 | 145 | Rep 3 | 147
Rep 4 | 145 | Rep 4 | 145 | Rep 4 | 7.3 | Rep 4 | 7.2 | Rep 4 | 145 | Rep 4 | 146
Rep 5 | 145 | Rep 5 | 144 | Rep 5 | 7.2 | Rep 5 | 7.3 | Rep 5 | 145 | Rep 5 | 145
Target | 145 | % Diff. | 0 | Target | 7.24 | % Diff. | 0.82873 | Target | 145.2 | % Diff. | 0.55096
(mean) |  | (Rep 1) |  | (mean) |  | (Rep 1) |  | (mean) |  | (Rep 1)

## CaC_Chol_CK_CO2_GluC
| Carrier | 19 |  |  | Carrier | 20 |  |  | Carrier | 21
Donor |  | Recipient |  | Donor |  | Recipient |  | Donor |  | Recipient
(Cup 4) | ACET | (Cup 5) | CaC | (Cup 4) | CaC | (Cup 5) | ACET | (Cup 4) | ACET | (Cup 5) | Chol
Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results
Rep 1 | 13.4 | Rep 1 | 13.4 | Rep 1 | 146 | Rep 1 | 146 | Rep 1 | 266 | Rep 1 | 265
Rep 2 | 13.4 | Rep 2 | 13.3 | Rep 2 | 145 | Rep 2 | 145 | Rep 2 | 266 | Rep 2 | 266
Rep 3 | 13.4 | Rep 3 | 13.3 | Rep 3 | 144 | Rep 3 | 145 | Rep 3 | 267 | Rep 3 | 264
Rep 4 | 13.4 | Rep 4 | 13.3 | Rep 4 | 145 | Rep 4 | 145 | Rep 4 | 268 | Rep 4 | 266
Rep 5 | 13.4 | Rep 5 | 13.4 | Rep 5 | 145 | Rep 5 | 145 | Rep 5 | 267 | Rep 5 | 266
Target | 13.4 | % Diff. | 0 | Target | 145 | % Diff. | 0.68966 | Target | 266.8 | % Diff. | -0.67466
(mean) |  | (Rep 1) |  | (mean) |  | (Rep 1) |  | (mean) |  | (Rep 1)
| Carrier | 22 |  |  | Carrier | 23 |  |  | Carrier | 24
Donor |  | Recipient |  | Donor |  | Recipient |  | Donor |  | Recipient
(Cup 4) | Chol | (Cup 5) | ACET | (Cup 4) | ACET | (Cup 5) | CK | (Cup 4) | CK | (Cup 5) | ACET
Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results
Rep 1 | 146 | Rep 1 | 146 | Rep 1 | 672 | Rep 1 | 665 | Rep 1 | 145 | Rep 1 | 147
Rep 2 | 145 | Rep 2 | 146 | Rep 2 | 669 | Rep 2 | 668 | Rep 2 | 145 | Rep 2 | 147
Rep 3 | 145 | Rep 3 | 146 | Rep 3 | 667 | Rep 3 | 670 | Rep 3 | 145 | Rep 3 | 147
Rep 4 | 145 | Rep 4 | 146 | Rep 4 | 664 | Rep 4 | 663 | Rep 4 | 144 | Rep 4 | 146
Rep 5 | 144 | Rep 5 | 145 | Rep 5 | 669 | Rep 5 | 667 | Rep 5 | 144 | Rep 5 | 146
Target | 145 | % Diff. | 0.68966 | Target | 668.2 | % Diff. | -0.4789 | Target | 144.6 | % Diff. | 1.6598
(mean) |  | (Rep 1) |  | (mean) |  | (Rep 1) |  | (mean) |  | (Rep 1)
| Carrier | 25 |  |  | Carrier | 26 |  |  | Carrier | 27
Donor |  | Recipient |  | Donor |  | Recipient |  | Donor |  | Recipient
(Cup 4) | ACET | (Cup 5) | CO2 | (Cup 4) | CO2 | (Cup 5) | ACET | (Cup 4) | ACET | (Cup 5) | GluC
Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results
Rep 1 | 24 | Rep 1 | 23 | Rep 1 | 145 | Rep 1 | 145 | Rep 1 | 395 | Rep 1 | 390
Rep 2 | 24 | Rep 2 | 23 | Rep 2 | 144 | Rep 2 | 145 | Rep 2 | 393 | Rep 2 | 391
Rep 3 | 24 | Rep 3 | 23 | Rep 3 | 145 | Rep 3 | 144 | Rep 3 | 389 | Rep 3 | 387
Rep 4 | 24 | Rep 4 | 24 | Rep 4 | 144 | Rep 4 | 144 | Rep 4 | 393 | Rep 4 | 392
Rep 5 | 24 | Rep 5 | 23 | Rep 5 | 143 | Rep 5 | 144 | Rep 5 | 400 | Rep 5 | 394
Target | 24 | % Diff. | -4.1667 | Target | 144.2 | % Diff. | 0.55479 | Target | 394 | % Diff. | -1.0152
(mean) |  | (Rep 1) |  | (mean) |  | (Rep 1) |  | (mean) |  | (Rep 1)

## GluC_Na-C_K-C_Cl-C_CrEnz
| Carrier | 28 |  |  | Carrier | 29 |  |  | Carrier | 30
Donor |  | Recipient |  | Donor |  | Recipient |  | Donor |  | Recipient
(Cup 4) | GluC | (Cup5) | ACET | (Cup 4) | ACET | (Cup 5) | Na-C | (Cup 4) | Na-C | (Cup 5) | ACET
Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results |  |  |  | Cup 5 Results
Rep 1 | 145 | Rep 1 | 147 | Rep 1 | 164 | Rep 1 | 164 | Rep 1 | 145 | Rep 1 | 144
Rep 2 | 145 | Rep 2 | 147 | Rep 2 | 164 | Rep 2 | 163 | Rep 2 | 144 | Rep 2 | 145
Rep 3 | 145 | Rep 3 | 146 | Rep 3 | 164 | Rep 3 | 164 | Rep 3 | 145 | Rep 3 | 145
Rep 4 | 144 | Rep 4 | 146 | Rep 4 | 164 | Rep 4 | 164 | Rep 4 | 145 | Rep 4 | 145
Rep 5 | 144 | Rep 5 | 145 | Rep 5 | 163 | Rep 5 | 163 | Rep 5 | 145 | Rep 5 | 144
Target | 144.6 | % Diff. | 1.6598 | Target | 163.8 | % Diff. | 0.1221 | Target | 144.8 | % Diff. | -0.55249
(mean) |  | (Rep 1) |  | (mean) |  | (Rep 1) |  | (mean) |  | (Rep 1)
| Carrier | 31 |  |  | Carrier | 32 |  |  | Carrier | 33
Donor |  | Recipient |  | Donor |  | Recipient |  | Donor |  | Recipient
(Cup 4) | ACET | (Cup 5) | Cl-C | (Cup 4) | Cl-C | (Cup 5) | ACET | (Cup 4) | ACET | (Cup 5) | K-C
Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results
Rep 1 | 125 | Rep 1 | 125 | Rep 1 | 145 | Rep 1 | 147 | Rep 1 | 7.8 | Rep 1 | 7.8
Rep 2 | 125 | Rep 2 | 125 | Rep 2 | 146 | Rep 2 | 147 | Rep 2 | 7.8 | Rep 2 | 7.8
Rep 3 | 125 | Rep 3 | 125 | Rep 3 | 145 | Rep 3 | 147 | Rep 3 | 7.8 | Rep 3 | 7.8
Rep 4 | 125 | Rep 4 | 125 | Rep 4 | 145 | Rep 4 | 147 | Rep 4 | 7.8 | Rep 4 | 7.8
Rep 5 | 123 | Rep 5 | 122 | Rep 5 | 145 | Rep 5 | 147 | Rep 5 | 7.8 | Rep 5 | 7.8
Target | 124.6 | % Diff. | 0.32103 | Target | 145.2 | % Diff. | 1.2397 | Target | 7.8 | % Diff. | 0
(mean) |  | (Rep 1) |  | (mean) |  | (Rep 1) |  | (mean) |  | (Rep 1)
| Carrier | 34 |  |  | Carrier | 35 |  |  | Carrier | 36
Donor |  | Recipient |  | Donor |  | Recipient |  | Donor |  | Recipient
(Cup 4) | K-C | (Cup 5) | ACET | (Cup 4) | ACET | (Cup 5) | CrEnz | (Cup 4) | CrEnz | (Cup 5) | ACET
Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results
Rep 1 | 145 | Rep 1 | 146 | Rep 1 | 7.14 | Rep 1 | 7.1 | Rep 1 | 145 | Rep 1 | 147
Rep 2 | 145 | Rep 2 | 146 | Rep 2 | 7.06 | Rep 2 | 7.09 | Rep 2 | 146 | Rep 2 | 147
Rep 3 | 145 | Rep 3 | 145 | Rep 3 | 7.1 | Rep 3 | 7.03 | Rep 3 | 146 | Rep 3 | 146
Rep 4 | 145 | Rep 4 | 145 | Rep 4 | 7.11 | Rep 4 | 7.05 | Rep 4 | 145 | Rep 4 | 147
Rep 5 | 145 | Rep 5 | 145 | Rep 5 | 7.11 | Rep 5 | 7.09 | Rep 5 | 145 | Rep 5 | 145
Target | 145 | % Diff. | 0.68966 | Target | 7.104 | % Diff. | -0.056306 | Target | 145.4 | % Diff. | 1.1004
(mean) |  | (Rep 1) |  | (mean) |  | (Rep 1) |  | (mean) |  | (Rep 1)

## Lact_Mg_Phos_TP_Urea
| Carrier | 37 |  |  | Carrier | 38 |  |  | Carrier | 39
Donor |  | Recipient |  | Donor |  | Recipient |  | Donor |  | Recipient
(Cup 4) | ACET | (Cup 5) | Lact | (Cup 4) | Lact | (Cup 5) | ACET | (Cup 4) | ACET | (Cup 5) | Mg
Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results
Rep 1 | 5.65 | Rep 1 | 5.67 | Rep 1 | 146 | Rep 1 | 146 | Rep 1 | 4.1 | Rep 1 | 4
Rep 2 | 5.67 | Rep 2 | 5.63 | Rep 2 | 146 | Rep 2 | 145 | Rep 2 | 4 | Rep 2 | 4
Rep 3 | 5.67 | Rep 3 | 5.68 | Rep 3 | 145 | Rep 3 | 145 | Rep 3 | 4 | Rep 3 | 4
Rep 4 | 5.69 | Rep 4 | 5.67 | Rep 4 | 145 | Rep 4 | 145 | Rep 4 | 4 | Rep 4 | 4
Rep 5 | 5.78 | Rep 5 | 5.76 | Rep 5 | 140 | Rep 5 | 144 | Rep 5 | 4 | Rep 5 | 3.9
Target | 5.692 | % Diff. | -0.38651 | Target | 144.4 | % Diff. | 1.108 | Target | 4.02 | % Diff. | -0.49751
(mean) |  | (Rep 1) |  | (mean) |  | (Rep 1) |  | (mean) |  | (Rep 1)
| Carrier | 40 |  |  | Carrier | 41 |  |  | Carrier | 42
Donor |  | Recipient |  | Donor |  | Recipient |  | Donor |  | Recipient
(Cup 4) | Mg | (Cup 5) | ACET | (Cup 4) | ACET | (Cup 5) | Phos | (Cup 4) | Phos | (Cup 5) | ACET
Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results
Rep 1 | 146 | Rep 1 | 147 | Rep 1 | 7.9 | Rep 1 | 7.9 | Rep 1 | 146 | Rep 1 | 146
Rep 2 | 145 | Rep 2 | 147 | Rep 2 | 7.9 | Rep 2 | 8 | Rep 2 | 145 | Rep 2 | 146
Rep 3 | 144 | Rep 3 | 146 | Rep 3 | 7.9 | Rep 3 | 7.9 | Rep 3 | 145 | Rep 3 | 145
Rep 4 | 145 | Rep 4 | 147 | Rep 4 | 7.9 | Rep 4 | 7.9 | Rep 4 | 145 | Rep 4 | 145
Rep 5 | 145 | Rep 5 | 146 | Rep 5 | 7.7 | Rep 5 | 7.9 | Rep 5 | 145 | Rep 5 | 146
Target | 145 | % Diff. | 1.3793 | Target | 7.86 | % Diff. | 0.50891 | Target | 145.2 | % Diff. | 0.55096
(mean) |  | (Rep 1) |  | (mean) |  | (Rep 1) |  | (mean) |  | (Rep 1)
| Carrier | 43 |  |  | Carrier | 44 |  |  | Carrier | 45
Donor |  | Recipient |  | Donor |  | Recipient |  | Donor |  | Recipient
(Cup 4) | ACET | (Cup 5) | TP | (Cup 4) | TP | (Cup 5) | ACET | (Cup 4) | ACET | (Cup 5) | Urea
Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results
Rep 1 | 6.7 | Rep 1 | 6.7 | Rep 1 | 146 | Rep 1 | 147 | Rep 1 | 72 | Rep 1 | 72
Rep 2 | 6.7 | Rep 2 | 6.7 | Rep 2 | 145 | Rep 2 | 147 | Rep 2 | 73 | Rep 2 | 73
Rep 3 | 6.7 | Rep 3 | 6.7 | Rep 3 | 145 | Rep 3 | 147 | Rep 3 | 72 | Rep 3 | 73
Rep 4 | 6.7 | Rep 4 | 6.7 | Rep 4 | 144 | Rep 4 | 147 | Rep 4 | 73 | Rep 4 | 73
Rep 5 | 6.7 | Rep 5 | 6.7 | Rep 5 | 142 | Rep 5 | 146 | Rep 5 | 74 | Rep 5 | 72
Target | 6.7 | % Diff. | 0 | Target | 144.4 | % Diff. | 1.8006 | Target | 72.8 | % Diff. | -1.0989
(mean) |  | (Rep 1) |  | (mean) |  | (Rep 1) |  | (mean) |  | (Rep 1)

## Urea_Trig_Sali_UA_UHDL
| Carrier | 46 |  |  | Carrier | 47 |  |  | Carrier | 48
Donor |  | Recipient |  | Donor |  | Recipient |  | Donor |  | Recipient
(Cup 4) | Urea | (Cup 5) | ACET | (Cup 4) | ACET | (Cup 5) | Trig | (Cup 4) | Trig | (Cup 5) | ACET
Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results
Rep 1 | 146 | Rep 1 | 146 | Rep 1 | 212 | Rep 1 | 214 | Rep 1 | 145 | Rep 1 | 146
Rep 2 | 146 | Rep 2 | 145 | Rep 2 | 211 | Rep 2 | 213 | Rep 2 | 145 | Rep 2 | 146
Rep 3 | 145 | Rep 3 | 145 | Rep 3 | 213 | Rep 3 | 214 | Rep 3 | 145 | Rep 3 | 146
Rep 4 | 145 | Rep 4 | 145 | Rep 4 | 215 | Rep 4 | 214 | Rep 4 | 144 | Rep 4 | 146
Rep 5 | 144 | Rep 5 | 144 | Rep 5 | 214 | Rep 5 | 214 | Rep 5 | 143 | Rep 5 | 146
Target | 145.2 | % Diff. | 0.55096 | Target | 213 | % Diff. | 0.46948 | Target | 144.4 | % Diff. | 1.108
(mean) |  | (Rep 1) |  | (mean) |  | (Rep 1) |  | (mean) |  | (Rep 1)
| Carrier | 49 |  |  | Carrier | 50 |  |  | Carrier | 51
Donor |  | Recipient |  | Donor |  | Recipient |  | Donor |  | Recipient
(Cup 4) | ACET | (Cup 5) | Sali | (Cup 4) | Sali | (Cup 5) | ACET | (Cup 4) | ACET | (Cup 5) | UA
Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results
Rep 1 | 19.4 | Rep 1 | 19.4 | Rep 1 | 144 | Rep 1 | 145 | Rep 1 | 9.4 | Rep 1 | 9.5
Rep 2 | 19.6 | Rep 2 | 19.3 | Rep 2 | 145 | Rep 2 | 145 | Rep 2 | 9.4 | Rep 2 | 9.4
Rep 3 | 19.3 | Rep 3 | 19.6 | Rep 3 | 144 | Rep 3 | 145 | Rep 3 | 9.4 | Rep 3 | 9.4
Rep 4 | 19.1 | Rep 4 | 19.7 | Rep 4 | 144 | Rep 4 | 144 | Rep 4 | 9.4 | Rep 4 | 9.4
Rep 5 | 18.6 | Rep 5 | 19.3 | Rep 5 | 143 | Rep 5 | 144 | Rep 5 | 9.4 | Rep 5 | 9.4
Target | 19.2 | % Diff. | 1.0417 | Target | 144 | % Diff. | 0.69444 | Target | 9.4 | % Diff. | 1.0638
(mean) |  | (Rep 1) |  | (mean) |  | (Rep 1) |  | (mean) |  | (Rep 1)
| Carrier | 52 |  |  | Carrier | 53 |  |  | Carrier | 54
Donor |  | Recipient |  | Donor |  | Recipient |  | Donor |  | Recipient
(Cup 4) | UA | (Cup 5) | ACET | (Cup 4) | ACET | (Cup 5) | UHDL | (Cup 4) | UHDL | (Cup 5) | ACET
Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results
Rep 1 | 146 | Rep 1 | 148 | Rep 1 | 75 | Rep 1 | 75 | Rep 1 | 146 | Rep 1 | 145
Rep 2 | 145 | Rep 2 | 147 | Rep 2 | 74 | Rep 2 | 75 | Rep 2 | 146 | Rep 2 | 145
Rep 3 | 145 | Rep 3 | 147 | Rep 3 | 75 | Rep 3 | 74 | Rep 3 | 145 | Rep 3 | 145
Rep 4 | 145 | Rep 4 | 147 | Rep 4 | 74 | Rep 4 | 74 | Rep 4 | 145 | Rep 4 | 145
Rep 5 | 145 | Rep 5 | 147 | Rep 5 | 75 | Rep 5 | 75 | Rep 5 | 144 | Rep 5 | 145
Target | 145.2 | % Diff. | 1.9284 | Target | 74.6 | % Diff. | 0.53619 | Target | 145.2 | % Diff. | -0.13774
(mean) |  | (Rep 1) |  | (mean) |  | (Rep 1) |  | (mean) |  | (Rep 1)

## Na-CU_K-CU_Cl-CU_AmpQ_BarbQ
| Carrier | 55 |  |  | Carrier | 56 |  |  | Carrier | 57
Donor |  | Recipient |  | Donor |  | Recipient |  | Donor |  | Recipient
(Cup 4) | ACET | (Cup 5) | Na-CU | (Cup 4) | Na-CU | (Cup 5) | ACET | (Cup 4) | ACET | (Cup 5) | K-CU
Cup 2 Results |  | Cup 5 Results |  |  |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results
Rep 1 | 170 | Rep 1 | 169 | Rep 1 | 143 | Rep 1 | 144 | Rep 1 | 69.3 | Rep 1 | 69.3
Rep 2 | 169 | Rep 2 | 169 | Rep 2 | 143 | Rep 2 | 144 | Rep 2 | 69.4 | Rep 2 | 69.1
Rep 3 | 169 | Rep 3 | 169 | Rep 3 | 143 | Rep 3 | 143 | Rep 3 | 69.2 | Rep 3 | 69.1
Rep 4 | 169 | Rep 4 | 168 | Rep 4 | 142 | Rep 4 | 143 | Rep 4 | 68.9 | Rep 4 | 68.7
Rep 5 | 168 | Rep 5 | 166 | Rep 5 | 141 | Rep 5 | 143 | Rep 5 | 67.8 | Rep 5 | 67.6
Target | 169 | % Diff. | 0 | Target | 142.4 | % Diff. | 1.1236 | Target | 68.92 | % Diff. | 0.55136
(mean) |  | (Rep 1) |  | (mean) |  | (Rep 1) |  | (mean) |  | (Rep 1)
| Carrier | 58 |  |  | Carrier | 59 |  |  | Carrier | 60
Donor |  | Recipient |  | Donor |  | Recipient |  | Donor |  | Recipient
(Cup 4) | K-CU | (Cup 5) | ACET | (Cup 4) | ACET | (Cup 5) | Cl-CU | (Cup 4) | Cl-CU | (Cup 5) | ACET
Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results
Rep 1 | 143 | Rep 1 | 143 | Rep 1 | 197 | Rep 1 | 197 | Rep 1 | 144 | Rep 1 | 143
Rep 2 | 143 | Rep 2 | 143 | Rep 2 | 197 | Rep 2 | 197 | Rep 2 | 144 | Rep 2 | 143
Rep 3 | 143 | Rep 3 | 144 | Rep 3 | 197 | Rep 3 | 196 | Rep 3 | 143 | Rep 3 | 143
Rep 4 | 142 | Rep 4 | 143 | Rep 4 | 195 | Rep 4 | 196 | Rep 4 | 143 | Rep 4 | 143
Rep 5 | 142 | Rep 5 | 144 | Rep 5 | 193 | Rep 5 | 193 | Rep 5 | 142 | Rep 5 | 143
Target | 142.6 | % Diff. | 0.2805 | Target | 195.8 | % Diff. | 0.61287 | Target | 143.2 | % Diff. | -0.13966
(mean) |  | (Rep 1) |  | (mean) |  | (Rep 1) |  | (mean) |  | (Rep 1)
| Carrier | 61 |  |  | Carrier | 62 |  |  | Carrier | 63
Donor |  | Recipient |  | Donor |  | Recipient |  | Donor |  | Recipient
(Cup 4) | ACET | (Cup 5) | AmpQ | (Cup 4) | AmpQ | (Cup 5) | ACET | (Cup 4) | ACET | (Cup 5) | BarbQ
Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results
Rep 1 | 1246 | Rep 1 | 1236 | Rep 1 | 143 | Rep 1 | 143 | Rep 1 | 318 | Rep 1 | 318
Rep 2 | 1227 | Rep 2 | 1212 | Rep 2 | 143 | Rep 2 | 142 | Rep 2 | 314 | Rep 2 | 310
Rep 3 | 1246 | Rep 3 | 1234 | Rep 3 | 142 | Rep 3 | 142 | Rep 3 | 319 | Rep 3 | 320
Rep 4 | 1240 | Rep 4 | 1233 | Rep 4 | 143 | Rep 4 | 142 | Rep 4 | 317 | Rep 4 | 319
Rep 5 | 1259 | Rep 5 | 1220 | Rep 5 | 143 | Rep 5 | 141 | Rep 5 | 316 | Rep 5 | 314
Target | 1243.6 | % Diff. | -0.61113 | Target | 142.8 | % Diff. | 0.14006 | Target | 316.8 | % Diff. | 0.37879
(mean) |  | (Rep 1) |  | (mean) |  | (Rep 1) |  | (mean) |  | (Rep 1)

## BarbQ_CocQ_MethQ_OpiQ_OXYCOD
| Carrier | 64 |  |  | Carrier | 65 |  |  | Carrier | 66
Donor |  | Recipient |  | Donor |  | Recipient |  | Donor |  | Recipient
(Cup 4) | BarbQ | (Cup 5) | ACET | (Cup 4) | ACET | (Cup 5) | CocQ | (Cup 4) | CocQ | (Cup 5) | ACET
Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results
Rep 1 | 143 | Rep 1 | 143 | Rep 1 | 357 | Rep 1 | 359 | Rep 1 | 143 | Rep 1 | 143
Rep 2 | 143 | Rep 2 | 143 | Rep 2 | 348 | Rep 2 | 354 | Rep 2 | 143 | Rep 2 | 144
Rep 3 | 143 | Rep 3 | 143 | Rep 3 | 352 | Rep 3 | 356 | Rep 3 | 143 | Rep 3 | 143
Rep 4 | 143 | Rep 4 | 142 | Rep 4 | 357 | Rep 4 | 355 | Rep 4 | 143 | Rep 4 | 142
Rep 5 | 143 | Rep 5 | 141 | Rep 5 | 355 | Rep 5 | 353 | Rep 5 | 143 | Rep 5 | 143
Target | 143 | % Diff. | 0 | Target | 353.8 | % Diff. | 1.4698 | Target | 143 | % Diff. | 0
(mean) |  | (Rep 1) |  | (mean) |  | (Rep 1) |  | (mean) |  | (Rep 1)
| Carrier | 67 |  |  | Carrier | 68 |  |  | Carrier | 69
Donor |  | Recipient |  | Donor |  | Recipient |  | Donor |  | Recipient
(Cup 4) | ACET | (Cup 5) | MethQ | (Cup 4) | MethQ | (Cup 5) | ACET | (Cup 4) | ACET | (Cup 5) | OpiQ
Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results
Rep 1 | 329 | Rep 1 | 327 | Rep 1 | 144 | Rep 1 | 143 | Rep 1 | 2307 | Rep 1 | 2304
Rep 2 | 326 | Rep 2 | 333 | Rep 2 | 144 | Rep 2 | 143 | Rep 2 | 2339 | Rep 2 | 2330
Rep 3 | 330 | Rep 3 | 322 | Rep 3 | 144 | Rep 3 | 143 | Rep 3 | 2332 | Rep 3 | 2354
Rep 4 | 328 | Rep 4 | 327 | Rep 4 | 143 | Rep 4 | 143 | Rep 4 | 2320 | Rep 4 | 2339
Rep 5 | 329 | Rep 5 | 322 | Rep 5 | 143 | Rep 5 | 140 | Rep 5 | 2313 | Rep 5 | 2313
Target | 328.4 | % Diff. | -0.42631 | Target | 143.6 | % Diff. | -0.41783 | Target | 2322.2 | % Diff. | -0.78374
(mean) |  | (Rep 1) |  | (mean) |  | (Rep 1) |  | (mean) |  | (Rep 1)
| Carrier | 70 |  |  | Carrier | 71 |  |  | Carrier | 72
Donor |  | Recipient |  | Donor |  | Recipient |  | Donor |  | Recipient
(Cup 4) | OpiQ | (Cup 5) | ACET | (Cup 4) | ACET | (Cup 5) | OXYCOD* | (Cup 4) | OXYCOD* | (Cup 5) | ACET
Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results
Rep 1 | 144 | Rep 1 | 144 | Rep 1 | 396.47 | Rep 1 | 393.48 | Rep 1 | 145 | Rep 1 | 145
Rep 2 | 144 | Rep 2 | 144 | Rep 2 | 397.47 | Rep 2 | 394.41 | Rep 2 | 145 | Rep 2 | 145
Rep 3 | 144 | Rep 3 | 144 | Rep 3 | 396.41 | Rep 3 | 393.08 | Rep 3 | 145 | Rep 3 | 144
Rep 4 | 143 | Rep 4 | 143 | Rep 4 | 396.21 | Rep 4 | 389.62 | Rep 4 | 145 | Rep 4 | 144
Rep 5 | 144 | Rep 5 | 142 | Rep 5 | 396.54 | Rep 5 | 394.81 | Rep 5 | 145 | Rep 5 | 142
Target | 143.8 | % Diff. | 0.13908 | Target | 396.62 | % Diff. | -0.79169 | Target | 145 | % Diff. | 0
(mean) |  | (Rep 1) |  | (mean) |  | (Rep 1) |  | (mean) |  | (Rep 1)
*Non-Abbott application

## EtohM_CRP32__Tobra_Vanco_LIDOCA
| Carrier | 73 |  |  | Carrier | 74 |  |  | Carrier | 75
Donor |  | Recipient |  | Donor |  | Recipient |  | Donor |  | Recipient
(Cup 4) | ACET | (Cup 5) | EtohM | (Cup 4) | EtohM | (Cup 5) | ACET | (Cup 4) | ACET | (Cup 5) | CRP32
Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results
Rep 1 | 95.5 | Rep 1 | 94.7 | Rep 1 | 145 | Rep 1 | 145 | Rep 1 | 6.34 | Rep 1 | 6.24
Rep 2 | 95.2 | Rep 2 | 95 | Rep 2 | 145 | Rep 2 | 145 | Rep 2 | 6.35 | Rep 2 | 6.39
Rep 3 | 95.6 | Rep 3 | 94.9 | Rep 3 | 144 | Rep 3 | 145 | Rep 3 | 6.29 | Rep 3 | 6.3
Rep 4 | 96 | Rep 4 | 95 | Rep 4 | 144 | Rep 4 | 144 | Rep 4 | 6.38 | Rep 4 | 6.3
Rep 5 | 95.2 | Rep 5 | 93.9 | Rep 5 | 142 | Rep 5 | 143 | Rep 5 | 6.47 | Rep 5 | 6.41
Target | 95.5 | % Diff. | -0.8377 | Target | 144 | % Diff. | 0.69444 | Target | 6.366 | % Diff. | -1.9793
(mean) |  | (Rep 1) |  | (mean) |  | (Rep 1) |  | (mean) |  | (Rep 1)
| Carrier | 76 |  |  | Carrier | 77 |  |  | Carrier | 78
Donor |  | Recipient |  | Donor |  | Recipient |  | Donor |  | Recipient
(Cup 4) | CRP32 | (Cup 5) | ACET | (Cup 4) | ACET | (Cup 5) | Tobra | (Cup 4) | Tobra | (Cup 5) | ACET
Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results
Rep 1 | 144 | Rep 1 | 143 | Rep 1 | 6.5 | Rep 1 | 6.5 | Rep 1 | 144 | Rep 1 | 144
Rep 2 | 144 | Rep 2 | 144 | Rep 2 | 6.5 | Rep 2 | 6.5 | Rep 2 | 143 | Rep 2 | 144
Rep 3 | 144 | Rep 3 | 143 | Rep 3 | 6.5 | Rep 3 | 6.5 | Rep 3 | 143 | Rep 3 | 144
Rep 4 | 143 | Rep 4 | 142 | Rep 4 | 6.4 | Rep 4 | 6.5 | Rep 4 | 143 | Rep 4 | 144
Rep 5 | 141 | Rep 5 | 141 | Rep 5 | 6.3 | Rep 5 | 6.3 | Rep 5 | 142 | Rep 5 | 144
Target | 143.2 | % Diff. | -0.13966 | Target | 6.44 | % Diff. | 0.93168 | Target | 143 | % Diff. | 0.6993
(mean) |  | (Rep 1) |  | (mean) |  | (Rep 1) |  | (mean) |  | (Rep 1)
| Carrier | 79 |  |  | Carrier | 80 |  |  | Carrier | 81
Donor |  | Recipient |  | Donor |  | Recipient |  | Donor |  | Recipient
(Cup 4) | ACET | (Cup 5) | Vanco | (Cup 4) | Vanco | (Cup 5) | ACET | (Cup 4) | ACET | (Cup 5) | LIDOCAIN*
Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results
Rep 1 | 34.8 | Rep 1 | 34.9 | Rep 1 | 144 | Rep 1 | 144 | Rep 1 | 7.16 | Rep 1 | 7.28
Rep 2 | 35 | Rep 2 | 34.7 | Rep 2 | 144 | Rep 2 | 143 | Rep 2 | 7.19 | Rep 2 | 7.04
Rep 3 | 34.9 | Rep 3 | 34.8 | Rep 3 | 144 | Rep 3 | 144 | Rep 3 | 7.19 | Rep 3 | 7.21
Rep 4 | 34.7 | Rep 4 | 34.9 | Rep 4 | 143 | Rep 4 | 143 | Rep 4 | 7.25 | Rep 4 | 7.24
Rep 5 | 35.1 | Rep 5 | 34.8 | Rep 5 | 143 | Rep 5 | 142 | Rep 5 | 7.23 | Rep 5 | 7.29
Target | 34.9 | % Diff. | 2.0359e-14 | Target | 143.6 | % Diff. | 0.27855 | Target | 7.204 | % Diff. | 1.055
(mean) |  | (Rep 1) |  | (mean) |  | (Rep 1) |  | (mean) |  | (Rep 1)
*Non-Abbott application

## LIDOCA_NAPA_PRIM_PROCA_TCASQ
| Carrier | 82 |  |  | Carrier | 83 |  |  | Carrier | 84
Donor |  | Recipient |  | Donor |  | Recipient |  | Donor |  | Recipient
(Cup 4) | LIDOCAIN* | (Cup 5) | ACET | (Cup 4) | ACET | (Cup 5) | NAPA | (Cup 4) | NAPA | (Cup 5) | ACET
Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results
Rep 1 | 143 | Rep 1 | 142 | Rep 1 | 8.53 | Rep 1 | 8.61 | Rep 1 | 142 | Rep 1 | 143
Rep 2 | 143 | Rep 2 | 143 | Rep 2 | 8.58 | Rep 2 | 8.58 | Rep 2 | 142 | Rep 2 | 142
Rep 3 | 142 | Rep 3 | 142 | Rep 3 | 8.29 | Rep 3 | 8.55 | Rep 3 | 142 | Rep 3 | 142
Rep 4 | 142 | Rep 4 | 142 | Rep 4 | 8.56 | Rep 4 | 8.5 | Rep 4 | 141 | Rep 4 | 142
Rep 5 | 141 | Rep 5 | 141 | Rep 5 | 8.65 | Rep 5 | 8.57 | Rep 5 | 141 | Rep 5 | 140
Target | 142.2 | % Diff. | -0.14065 | Target | 8.522 | % Diff. | 1.0326 | Target | 141.6 | % Diff. | 0.9887
(mean) |  | (Rep 1) |  | (mean) |  | (Rep 1) |  | (mean) |  | (Rep 1)
| Carrier | 85 |  |  | Carrier | 86 |  |  | Carrier | 87
Donor |  | Recipient |  | Donor |  | Recipient |  | Donor |  | Recipient
(Cup 4) | ACET | (Cup 5) | PRIMIDON* | (Cup 4) | PRIMIDON* | (Cup 5) | ACET | (Cup 4) | ACET | (Cup 5) | PROCAIN*
Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results
Rep 1 | 11.24 | Rep 1 | 11.5 | Rep 1 | 143 | Rep 1 | 142 | Rep 1 | 12.18 | Rep 1 | 12.26
Rep 2 | 11.52 | Rep 2 | 11.65 | Rep 2 | 143 | Rep 2 | 143 | Rep 2 | 12.09 | Rep 2 | 12.08
Rep 3 | 11.34 | Rep 3 | 11.66 | Rep 3 | 143 | Rep 3 | 142 | Rep 3 | 11.76 | Rep 3 | 11.92
Rep 4 | 11.37 | Rep 4 | 11.18 | Rep 4 | 142 | Rep 4 | 142 | Rep 4 | 12.12 | Rep 4 | 12.26
Rep 5 | 11.04 | Rep 5 | 11.11 | Rep 5 | 142 | Rep 5 | 142 | Rep 5 | 11.97 | Rep 5 | 12.17
Target | 11.302 | % Diff. | 1.7519 | Target | 142.6 | % Diff. | -0.42076 | Target | 12.024 | % Diff. | 1.9627
(mean) |  | (Rep 1) |  | (mean) |  | (Rep 1) |  | (mean) |  | (Rep 1)
| Carrier | 88 |  |  | Carrier | 89 |  |  | Carrier | 90
Donor |  | Recipient |  | Donor |  | Recipient |  | Donor |  | Recipient
(Cup 4) | PROCAIN* | (Cup 5) | ACET | (Cup 4) | ACET | (Cup 5) | TCASQ | (Cup 4) | TCASQ | (Cup 5) | ACET
Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results
Rep 1 | 143 | Rep 1 | 143 | Rep 1 | 446 | Rep 1 | 434.5 | Rep 1 | 144 | Rep 1 | 144
Rep 2 | 143 | Rep 2 | 143 | Rep 2 | 443.1 | Rep 2 | 459.4 | Rep 2 | 144 | Rep 2 | 143
Rep 3 | 143 | Rep 3 | 143 | Rep 3 | 414.1 | Rep 3 | 429.9 | Rep 3 | 144 | Rep 3 | 144
Rep 4 | 143 | Rep 4 | 142 | Rep 4 | 454.9 | Rep 4 | 439.2 | Rep 4 | 143 | Rep 4 | 143
Rep 5 | 142 | Rep 5 | 140 | Rep 5 | 455.7 | Rep 5 | 443.7 | Rep 5 | 142 | Rep 5 | 143
Target | 142.8 | % Diff. | 0.14006 | Target | 442.76 | % Diff. | -1.8656 | Target | 143.4 | % Diff. | 0.41841
(mean) |  | (Rep 1) |  | (mean) |  | (Rep 1) |  | (mean) |  | (Rep 1)
*Non-Abbott application

## SBenSQ_Amm_Urea-U_Lip_Cerulo
| Carrier | 90 |  |  | Carrier | 91 |  |  | Carrier | 92
Donor |  | Recipient |  | Donor |  | Recipient |  | Donor |  | Recipient
(Cup 4) | ACET | (Cup 5) | SBenSQ | (Cup 4) | SBenSQ | (Cup 5) | ACET | (Cup 4) | ACET | (Cup 5) | Amm
Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results
Rep 1 | 93.7 | Rep 1 | 94.4 | Rep 1 | 144 | Rep 1 | 145 | Rep 1 | 294.98 | Rep 1 | 298.84
Rep 2 | 93.9 | Rep 2 | 94.1 | Rep 2 | 144 | Rep 2 | 144 | Rep 2 | 294.09 | Rep 2 | 298.7
Rep 3 | 94.2 | Rep 3 | 93.5 | Rep 3 | 144 | Rep 3 | 144 | Rep 3 | 289.49 | Rep 3 | 294.69
Rep 4 | 93.8 | Rep 4 | 93.2 | Rep 4 | 144 | Rep 4 | 145 | Rep 4 | 293.2 | Rep 4 | 299.88
Rep 5 | 99.1 | Rep 5 | 95.8 | Rep 5 | 142 | Rep 5 | 143 | Rep 5 | 294.98 | Rep 5 | 298.25
Target | 94.94 | % Diff. | -0.56878 | Target | 143.6 | % Diff. | 0.97493 | Target | 293.348 | % Diff. | 1.8722
(mean) |  | (Rep 1) |  | (mean) |  | (Rep 1) |  | (mean) |  | (Rep 1)
| Carrier | 93 |  |  | Carrier | 94 |  |  | Carrier | 95
Donor |  | Recipient |  | Donor |  | Recipient |  | Donor |  | Recipient
(Cup 4) | Amm | (Cup 5) | ACET | (Cup 4) | ACET | (Cup 5) | Urea-U | (Cup 4) | Urea-U | (Cup 5) | ACET
Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results
Rep 1 | 145 | Rep 1 | 145 | Rep 1 | 675 | Rep 1 | 673 | Rep 1 | 145 | Rep 1 | 144
Rep 2 | 144 | Rep 2 | 145 | Rep 2 | 685 | Rep 2 | 666 | Rep 2 | 144 | Rep 2 | 144
Rep 3 | 144 | Rep 3 | 144 | Rep 3 | 679 | Rep 3 | 681 | Rep 3 | 144 | Rep 3 | 144
Rep 4 | 144 | Rep 4 | 145 | Rep 4 | 693 | Rep 4 | 686 | Rep 4 | 144 | Rep 4 | 145
Rep 5 | 144 | Rep 5 | 143 | Rep 5 | 692 | Rep 5 | 705 | Rep 5 | 144 | Rep 5 | 144
Target | 144.2 | % Diff. | 0.55479 | Target | 684.8 | % Diff. | -1.7231 | Target | 144.2 | % Diff. | -0.1387
(mean) |  | (Rep 1) |  | (mean) |  | (Rep 1) |  | (mean) |  | (Rep 1)
| Carrier | 96 |  |  | Carrier | 97 |  |  | Carrier | 98
Donor |  | Recipient |  | Donor |  | Recipient |  | Donor |  | Recipient
(Cup 4) | ACET | (Cup 5) | Lip | (Cup 4) | Lip | (Cup 5) | ACET | (Cup 4) | ACET | (Cup 5) | Cerulo
Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results
Rep 1 | 158 | Rep 1 | 159 | Rep 1 | 147 | Rep 1 | 147 | Rep 1 | 42 | Rep 1 | 42
Rep 2 | 162 | Rep 2 | 159 | Rep 2 | 147 | Rep 2 | 147 | Rep 2 | 41 | Rep 2 | 41
Rep 3 | 163 | Rep 3 | 158 | Rep 3 | 146 | Rep 3 | 147 | Rep 3 | 41 | Rep 3 | 41
Rep 4 | 157 | Rep 4 | 162 | Rep 4 | 146 | Rep 4 | 148 | Rep 4 | 41 | Rep 4 | 41
Rep 5 | 154 | Rep 5 | 163 | Rep 5 | 146 | Rep 5 | 147 | Rep 5 | 42 | Rep 5 | 41
Target | 158.8 | % Diff. | 0.12594 | Target | 146.4 | % Diff. | 0.40984 | Target | 41.4 | % Diff. | 1.4493
(mean) |  | (Rep 1) |  | (mean) |  | (Rep 1) |  | (mean) |  | (Rep 1)

## Cerulo_C3_ASO_C4_LDH
| Carrier | 99 |  |  | Carrier | 100 |  |  | Carrier | 101
Donor |  | Recipient |  | Donor |  | Recipient |  | Donor |  | Recipient
(Cup 4) | Cerulo | (Cup 5) | ACET | (Cup 4) | ACET | (Cup 5) | C3 | (Cup 4) | C3 | (Cup 5) | ACET
Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results
Rep 1 | 142.16 | Rep 1 | 142.73 | Rep 1 | 244 | Rep 1 | 243 | Rep 1 | 145.1 | Rep 1 | 144.8
Rep 2 | 142.12 | Rep 2 | 142.55 | Rep 2 | 241 | Rep 2 | 239 | Rep 2 | 144.8 | Rep 2 | 144.6
Rep 3 | 141.88 | Rep 3 | 142.58 | Rep 3 | 240 | Rep 3 | 241 | Rep 3 | 144.5 | Rep 3 | 144.6
Rep 4 | 141.52 | Rep 4 | 142.18 | Rep 4 | 244 | Rep 4 | 239 | Rep 4 | 144 | Rep 4 | 144.4
Rep 5 | 141.27 | Rep 5 | 142.18 | Rep 5 | 246 | Rep 5 | 238 | Rep 5 | 144.1 | Rep 5 | 144
Target | 141.79 | % Diff. | 0.66295 | Target | 243 | % Diff. | 0 | Target | 144.5 | % Diff. | 0.20761
(mean) |  | (Rep 1) |  | (mean) |  | (Rep 1) |  | (mean) |  | (Rep 1)
| Carrier | 102 |  |  | Carrier | 103 |  |  | Carrier | 104
Donor |  | Recipient |  | Donor |  | Recipient |  | Donor |  | Recipient
(Cup 4) | ACET | (Cup 5) | ASO | (Cup 4) | ASO | (Cup 5) | ACET | (Cup 4) | ACET | (Cup 5) | C4
Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results
Rep 1 | 123.7 | Rep 1 | 123.6 | Rep 1 | 145.1 | Rep 1 | 145.2 | Rep 1 | 47.4 | Rep 1 | 47.6
Rep 2 | 124.8 | Rep 2 | 122.3 | Rep 2 | 144.1 | Rep 2 | 144.9 | Rep 2 | 47.6 | Rep 2 | 47.6
Rep 3 | 122.8 | Rep 3 | 124 | Rep 3 | 144.7 | Rep 3 | 144.1 | Rep 3 | 48.6 | Rep 3 | 47.7
Rep 4 | 124.1 | Rep 4 | 123.2 | Rep 4 | 144 | Rep 4 | 144 | Rep 4 | 48 | Rep 4 | 47.2
Rep 5 | 123.2 | Rep 5 | 124.1 | Rep 5 | 143.2 | Rep 5 | 143.9 | Rep 5 | 47.8 | Rep 5 | 47.3
Target | 123.72 | % Diff. | -0.096993 | Target | 144.22 | % Diff. | 0.67952 | Target | 47.88 | % Diff. | -0.5848
(mean) |  | (Rep 1) |  | (mean) |  | (Rep 1) |  | (mean) |  | (Rep 1)
| Carrier | 105 |  |  | Carrier | 106 |  |  | Carrier | 107
Donor |  | Recipient |  | Donor |  | Recipient |  | Donor |  | Recipient
(Cup 4) | C4 | (Cup 5) | ACET | (Cup 4) | ACET | (Cup 5) | LDH | (Cup 4) | LDH | (Cup 5) | ACET
Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results
Rep 1 | 145.8 | Rep 1 | 145.4 | Rep 1 | 401 | Rep 1 | 397 | Rep 1 | 145 | Rep 1 | 144
Rep 2 | 144.9 | Rep 2 | 145.4 | Rep 2 | 399 | Rep 2 | 389 | Rep 2 | 144 | Rep 2 | 144
Rep 3 | 145.3 | Rep 3 | 144.9 | Rep 3 | 392 | Rep 3 | 390 | Rep 3 | 145 | Rep 3 | 144
Rep 4 | 144.7 | Rep 4 | 144.6 | Rep 4 | 402 | Rep 4 | 394 | Rep 4 | 144 | Rep 4 | 144
Rep 5 | 145 | Rep 5 | 144.3 | Rep 5 | 400 | Rep 5 | 393 | Rep 5 | 144 | Rep 5 | 144
Target | 145.14 | % Diff. | 0.17914 | Target | 398.8 | % Diff. | -0.45135 | Target | 144.4 | % Diff. | -0.27701
(mean) |  | (Rep 1) |  | (mean) |  | (Rep 1) |  | (mean) |  | (Rep 1)

## Lip_TRF_Fe-Pl_GGT_Hapt
| Carrier | 108 |  |  | Carrier | 109 |  |  | Carrier | 110
Donor |  | Recipient |  | Donor |  | Recipient |  | Donor |  | Recipient
(Cup 4) | ACET | (Cup 5) | Lipase | (Cup 4) | Lipase | (Cup 5) | ACET | (Cup 4) | ACET | (Cup 5) | TRF
Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results
Rep 1 | 150 | Rep 1 | 147 | Rep 1 | 149.6 | Rep 1 | 150.2 | Rep 1 | 268 | Rep 1 | 268
Rep 2 | 148 | Rep 2 | 147 | Rep 2 | 149.4 | Rep 2 | 149.8 | Rep 2 | 270 | Rep 2 | 267
Rep 3 | 146 | Rep 3 | 146 | Rep 3 | 149.6 | Rep 3 | 149.8 | Rep 3 | 270 | Rep 3 | 270
Rep 4 | 146 | Rep 4 | 149 | Rep 4 | 149.9 | Rep 4 | 149.7 | Rep 4 | 269 | Rep 4 | 269
Rep 5 | 146 | Rep 5 | 146 | Rep 5 | 149.8 | Rep 5 | 150.3 | Rep 5 | 270 | Rep 5 | 270
Target | 147.2 | % Diff. | -0.13587 | Target | 149.66 | % Diff. | 0.36082 | Target | 269.4 | % Diff. | -0.51967
(mean) |  | (Rep 1) |  | (mean) |  | (Rep 1) |  | (mean) |  | (Rep 1)
| Carrier | 111 |  |  | Carrier | 112 |  |  | Carrier | 113
Donor |  | Recipient |  | Donor |  | Recipient |  | Donor |  | Recipient
(Cup 4) | TRF | (Cup 5) | ACET | (Cup 4) | ACET | (Cup 5) | Fe-PL | (Cup 4) | Fe-Pl | (Cup 5) | ACET
Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results
Rep 1 | 149.8 | Rep 1 | 149.5 | Rep 1 | 220 | Rep 1 | 220 | Rep 1 | 149 | Rep 1 | 149.3
Rep 2 | 148.9 | Rep 2 | 149.2 | Rep 2 | 224 | Rep 2 | 220 | Rep 2 | 148.5 | Rep 2 | 149
Rep 3 | 149.1 | Rep 3 | 149.9 | Rep 3 | 222 | Rep 3 | 222 | Rep 3 | 149 | Rep 3 | 148.2
Rep 4 | 148.7 | Rep 4 | 149.4 | Rep 4 | 222 | Rep 4 | 220 | Rep 4 | 147.9 | Rep 4 | 147.6
Rep 5 | 149.4 | Rep 5 | 150 | Rep 5 | 225 | Rep 5 | 222 | Rep 5 | 147.9 | Rep 5 | 149.3
Target | 149.18 | % Diff. | 0.21451 | Target | 222.6 | % Diff. | -1.168 | Target | 148.46 | % Diff. | 0.56581
(mean) |  | (Rep 1) |  | (mean) |  | (Rep 1) |  | (mean) |  | (Rep 1)
| Carrier | 114 |  |  | Carrier | 115 |  |  | Carrier | 116
Donor |  | Recipient |  | Donor |  | Recipient |  | Donor |  | Recipient
(Cup 4) | ACET | (Cup 5) | GGT | (Cup 4) | GGT | (Cup 5) | ACET | (Cup 4) | ACET | (Cup 5) | Hapt
Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results
Rep 1 | 139 | Rep 1 | 138 | Rep 1 | 149.4 | Rep 1 | 149.5 | Rep 1 | 226 | Rep 1 | 224
Rep 2 | 139 | Rep 2 | 139 | Rep 2 | 149.8 | Rep 2 | 150 | Rep 2 | 224 | Rep 2 | 222
Rep 3 | 139 | Rep 3 | 139 | Rep 3 | 149.6 | Rep 3 | 149.2 | Rep 3 | 223 | Rep 3 | 222
Rep 4 | 139 | Rep 4 | 137 | Rep 4 | 149.6 | Rep 4 | 149 | Rep 4 | 223 | Rep 4 | 221
Rep 5 | 140 | Rep 5 | 138 | Rep 5 | 150 | Rep 5 | 149.9 | Rep 5 | 225 | Rep 5 | 222
Target | 139.2 | % Diff. | -0.86207 | Target | 149.68 | % Diff. | -0.12026 | Target | 224.2 | % Diff. | -0.089206
(mean) |  | (Rep 1) |  | (mean) |  | (Rep 1) |  | (mean) |  | (Rep 1)

## Hapt_IgA_IgG_IgM_PAlb
| Carrier | 116 |  |  | Carrier | 117 |  |  | Carrier | 118
Donor |  | Recipient |  | Donor |  | Recipient |  | Donor |  | Recipient
(Cup 4) | Hapt | (Cup 5) | ACET | (Cup 4) | ACET | (Cup 5) | IgA | (Cup 4) | IgA | (Cup 5) | ACET
Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results
Rep 1 | 145.1 | Rep 1 | 145.3 | Rep 1 | 343 | Rep 1 | 345 | Rep 1 | 145.4 | Rep 1 | 144.9
Rep 2 | 145.5 | Rep 2 | 145.3 | Rep 2 | 348 | Rep 2 | 346 | Rep 2 | 145.3 | Rep 2 | 144.5
Rep 3 | 145.2 | Rep 3 | 145.3 | Rep 3 | 348 | Rep 3 | 347 | Rep 3 | 143.7 | Rep 3 | 144.2
Rep 4 | 144.2 | Rep 4 | 144.4 | Rep 4 | 347 | Rep 4 | 345 | Rep 4 | 143.6 | Rep 4 | 143.8
Rep 5 | 144.9 | Rep 5 | 143.9 | Rep 5 | 347 | Rep 5 | 347 | Rep 5 | 143.5 | Rep 5 | 144.1
Target | 144.98 | % Diff. | 0.22072 | Target | 346.6 | % Diff. | -0.46163 | Target | 144.3 | % Diff. | 0.4158
(mean) |  | (Rep 1) |  | (mean) |  | (Rep 1) |  | (mean) |  | (Rep 1)
| Carrier | 119 |  |  | Carrier | 120 |  |  | Carrier | 121
Donor |  | Recipient |  | Donor |  | Recipient |  | Donor |  | Recipient
(Cup 4) | ACET | (Cup 5) | IgG | (Cup 4) | IgG | (Cup 5) | ACET | (Cup 4) | ACET | (Cup 5) | IgM
Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results
Rep 1 | 1837 | Rep 1 | 1808 | Rep 1 | 145.3 | Rep 1 | 145.3 | Rep 1 | 182 | Rep 1 | 181
Rep 2 | 1794 | Rep 2 | 1809 | Rep 2 | 146 | Rep 2 | 145.3 | Rep 2 | 181 | Rep 2 | 182
Rep 3 | 1891 | Rep 3 | 1810 | Rep 3 | 145.4 | Rep 3 | 145.1 | Rep 3 | 182 | Rep 3 | 183
Rep 4 | 1845 | Rep 4 | 1845 | Rep 4 | 144.8 | Rep 4 | 143.6 | Rep 4 | 179 | Rep 4 | 182
Rep 5 | 1848 | Rep 5 | 1820 | Rep 5 | 144.4 | Rep 5 | 143.9 | Rep 5 | 182 | Rep 5 | 182
Target | 1843 | % Diff. | -1.8991 | Target | 145.18 | % Diff. | 0.082656 | Target | 181.2 | % Diff. | -0.11038
(mean) |  | (Rep 1) |  | (mean) |  | (Rep 1) |  | (mean) |  | (Rep 1)
| Carrier | 122 |  |  | Carrier | 123 |  |  | Carrier | 124
Donor |  | Recipient |  | Donor |  | Recipient |  | Donor |  | Recipient
(Cup 4) | IgM | (Cup 5) | ACET | (Cup 4) | ACET | (Cup 5) | PAlb | (Cup 4) | PAlb | (Cup 5) | ACET
Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results
Rep 1 | 145.9 | Rep 1 | 145.5 | Rep 1 | 28 | Rep 1 | 29 | Rep 1 | 146.3 | Rep 1 | 146.4
Rep 2 | 147.1 | Rep 2 | 145.3 | Rep 2 | 28 | Rep 2 | 28 | Rep 2 | 145.8 | Rep 2 | 146.5
Rep 3 | 145.4 | Rep 3 | 144.3 | Rep 3 | 28 | Rep 3 | 28 | Rep 3 | 145.8 | Rep 3 | 146
Rep 4 | 145.1 | Rep 4 | 145.1 | Rep 4 | 28 | Rep 4 | 28 | Rep 4 | 144 | Rep 4 | 145.5
Rep 5 | 144.3 | Rep 5 | 145.3 | Rep 5 | 29 | Rep 5 | 28 | Rep 5 | 144.5 | Rep 5 | 145
Target | 145.56 | % Diff. | -0.04122 | Target | 28.2 | % Diff. | 2.8369 | Target | 145.28 | % Diff. | 0.77093
(mean) |  | (Rep 1) |  | (mean) |  | (Rep 1) |  | (mean) |  | (Rep 1)

## RF_dLDL_Dig_Gent_VPA
| Carrier | 125 |  |  | Carrier | 126 |  |  | Carrier | 127
Donor |  | Recipient |  | Donor |  | Recipient |  | Donor |  | Recipient
(Cup 4) | ACET | (Cup 5) | RF | (Cup 4) | RF | (Cup 5) | ACET | (Cup 4) | ACET | (Cup 5) | DLDL
List Number | 3R11 |  | 8G66 | List Number | 8G66 |  | 3R11
Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results
Rep 1 | 49.8 | Rep 1 | 49.6 | Rep 1 | 145.9 | Rep 1 | 146.6 | Rep 1 | 141 | Rep 1 | 139
Rep 2 | 49.6 | Rep 2 | 49.3 | Rep 2 | 145.8 | Rep 2 | 146.7 | Rep 2 | 139 | Rep 2 | 138
Rep 3 | 49.7 | Rep 3 | 48.9 | Rep 3 | 145.6 | Rep 3 | 145.4 | Rep 3 | 140 | Rep 3 | 140
Rep 4 | 50.7 | Rep 4 | 49.8 | Rep 4 | 143.7 | Rep 4 | 145.5 | Rep 4 | 142 | Rep 4 | 140
Rep 5 | 49.8 | Rep 5 | 48.8 | Rep 5 | 144.6 | Rep 5 | 145.4 | Rep 5 | 140 | Rep 5 | 138
Target | 49.92 | % Diff. | -0.64103 | Target | 145.12 | % Diff. | 1.0198 | Target | 140.4 | % Diff. | -0.99715
(mean) |  | (Rep 1) |  | (mean) |  | (Rep 1) |  | (mean) |  | (Rep 1)
| Carrier | 128 |  |  | Carrier | 129 |  |  | Carrier | 130
Donor |  | Recipient |  | Donor |  | Recipient |  | Donor |  | Recipient
(Cup 4) | DLDL | (Cup 5) | ACET | (Cup 4) | ACET | (Cup 5) | Dig | (Cup 4) | Dig | (Cup 5) | ACET
Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results
Rep 1 | 146.1 | Rep 1 | 145.6 | Rep 1 | 3.24 | Rep 1 | 3.27 | Rep 1 | 145.4 | Rep 1 | 146.3
Rep 2 | 145.7 | Rep 2 | 145 | Rep 2 | 3.26 | Rep 2 | 3.26 | Rep 2 | 142.9 | Rep 2 | 145.5
Rep 3 | 144.7 | Rep 3 | 145.4 | Rep 3 | 3.24 | Rep 3 | 3.22 | Rep 3 | 144.7 | Rep 3 | 145.5
Rep 4 | 145.2 | Rep 4 | 145.1 | Rep 4 | 3.21 | Rep 4 | 3.21 | Rep 4 | 144 | Rep 4 | 144.9
Rep 5 | 145.7 | Rep 5 | 144.6 | Rep 5 | 3.16 | Rep 5 | 3.22 | Rep 5 | 143.2 | Rep 5 | 144.7
Target | 145.48 | % Diff. | 0.082486 | Target | 3.222 | % Diff. | 1.4898 | Target | 144.04 | % Diff. | 1.569
(mean) |  | (Rep 1) |  | (mean) |  | (Rep 1) |  | (mean) |  | (Rep 1)
| Carrier | 131 |  |  | Carrier | 132 |  |  | Carrier | 133
Donor |  | Recipient |  | Donor |  | Recipient |  | Donor |  | Recipient
(Cup 4) | ACET | (Cup 5) | Gent | (Cup 4) | Gent | (Cup 5) | ACET | (Cup 4) | ACET | (Cup 5) | VPA
Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results
Rep 1 | 9.1 | Rep 1 | 9 | Rep 1 | 147.2 | Rep 1 | 147.3 | Rep 1 | 119.8 | Rep 1 | 119.3
Rep 2 | 8.9 | Rep 2 | 9 | Rep 2 | 146.4 | Rep 2 | 147.1 | Rep 2 | 118.8 | Rep 2 | 119.6
Rep 3 | 8.9 | Rep 3 | 8.9 | Rep 3 | 146.7 | Rep 3 | 146.2 | Rep 3 | 118.8 | Rep 3 | 120
Rep 4 | 8.8 | Rep 4 | 8.8 | Rep 4 | 146.8 | Rep 4 | 146.4 | Rep 4 | 118.5 | Rep 4 | 118.7
Rep 5 | 8.8 | Rep 5 | 8.8 | Rep 5 | 145.6 | Rep 5 | 145.8 | Rep 5 | 118.3 | Rep 5 | 118.7
Target | 8.9 | % Diff. | 1.1236 | Target | 146.54 | % Diff. | 0.51863 | Target | 118.84 | % Diff. | 0.38708
(mean) |  | (Rep 1) |  | (mean) |  | (Rep 1) |  | (mean) |  | (Rep 1)

## VPA_uAlb_Upro_Benz_THC_Lith
| Carrier | 134 |  |  | Carrier | 135 |  |  | Carrier | 136
Donor |  | Recipient |  | Donor |  | Recipient |  | Donor |  | Recipient
(Cup 4) | VPA | (Cup 5) | ACET | (Cup 4) | ACET | (Cup 5) | uAlb | (Cup 4) | uAlb | (Cup 5) | ACET
Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results
Rep 1 | 147 | Rep 1 | 146 | Rep 1 | 9.2 | Rep 1 | 9.2 | Rep 1 | 145.2 | Rep 1 | 143.9
Rep 2 | 145.5 | Rep 2 | 146.4 | Rep 2 | 9.2 | Rep 2 | 9.1 | Rep 2 | 144.6 | Rep 2 | 144.7
Rep 3 | 145.4 | Rep 3 | 146.1 | Rep 3 | 9.3 | Rep 3 | 9.3 | Rep 3 | 144.6 | Rep 3 | 145
Rep 4 | 145 | Rep 4 | 145.4 | Rep 4 | 9.1 | Rep 4 | 9.2 | Rep 4 | 144.1 | Rep 4 | 145
Rep 5 | 144.6 | Rep 5 | 143.6 | Rep 5 | 9.1 | Rep 5 | 9.1 | Rep 5 | 143.8 | Rep 5 | 144.4
Target | 145.5 | % Diff. | 0.34364 | Target | 9.18 | % Diff. | 0.21786 | Target | 144.46 | % Diff. | -0.38765
(mean) |  | (Rep 1) |  | (mean) |  | (Rep 1) |  | (mean) |  | (Rep 1)
| Carrier | 137 |  |  | Carrier | 138 |  |  | Carrier | 139
Donor |  | Recipient |  | Donor |  | Recipient |  | Donor |  | Recipient
(Cup 4) | ACET | (Cup 5) | Upro | (Cup 4) | Upro | (Cup 5) | ACET | (Cup 4) | ACET | (Cup 5) | Benz
Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results
Rep 1 | 59 | Rep 1 | 59.7 | Rep 1 | 143.5 | Rep 1 | 143.8 | Rep 1 | 399.7 | Rep 1 | 392.3
Rep 2 | 59.7 | Rep 2 | 59.6 | Rep 2 | 143.6 | Rep 2 | 143.5 | Rep 2 | 401.3 | Rep 2 | 387.7
Rep 3 | 58.6 | Rep 3 | 58.5 | Rep 3 | 143.5 | Rep 3 | 143.6 | Rep 3 | 402.4 | Rep 3 | 395.4
Rep 4 | 58.7 | Rep 4 | 58.1 | Rep 4 | 142.7 | Rep 4 | 143.5 | Rep 4 | 390.6 | Rep 4 | 387.2
Rep 5 | 58.9 | Rep 5 | 58.5 | Rep 5 | 142.5 | Rep 5 | 142.8 | Rep 5 | 389.5 | Rep 5 | 392.4
Target | 58.98 | % Diff. | 1.2208 | Target | 143.16 | % Diff. | 0.44705 | Target | 396.7 | % Diff. | -1.1092
(mean) |  | (Rep 1) |  | (mean) |  | (Rep 1) |  | (mean) |  | (Rep 1)
| Carrier | 140 |  |  | Carrier | 141 |  |  | Carrier | 142
Donor |  | Recipient |  | Donor |  | Recipient |  | Donor |  | Recipient
(Cup 4) | ACET | (Cup 5) | THC | (Cup 4) | THC | (Cup 5) | ACET | (Cup 4) | ACET | (Cup 5) | Lith
Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results
Rep 1 | 56 | Rep 1 | 56.7 | Rep 1 | 143.9 | Rep 1 | 144 | Rep 1 | 2.283 | Rep 1 | 2.274
Rep 2 | 56.6 | Rep 2 | 56.5 | Rep 2 | 143.5 | Rep 2 | 143.8 | Rep 2 | 2.281 | Rep 2 | 2.303
Rep 3 | 58.6 | Rep 3 | 53.9 | Rep 3 | 143.9 | Rep 3 | 143.8 | Rep 3 | 2.267 | Rep 3 | 2.284
Rep 4 | 59.2 | Rep 4 | 56.5 | Rep 4 | 142.9 | Rep 4 | 142.7 | Rep 4 | 2.291 | Rep 4 | 2.298
Rep 5 | 55.3 | Rep 5 | 52.8 | Rep 5 | 142.4 | Rep 5 | 141.8 | Rep 5 | 2.301 | Rep 5 | 2.299
Target | 57.14 | % Diff. | -0.77004 | Target | 143.32 | % Diff. | 0.47446 | Target | 2.2846 | % Diff. | -0.46398
(mean) |  | (Rep 1) |  | (mean) |  | (Rep 1) |  | (mean) |  | (Rep 1)

## Lith_Amik_DIG_Quin_CreaC
| Carrier | 143 |  |  | Carrier | 144 |  |  | Carrier | 145
Donor |  | Recipient |  | Donor |  | Recipient |  | Donor |  | Recipient
(Cup 4) | Lith | (Cup 5) | ACET | (Cup 4) | ACET | (Cup 5) | Amik | (Cup 4) | Amik | (Cup 5) | ACET
Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results
Rep 1 | 146 | Rep 1 | 146 | Rep 1 | 29.1 | Rep 1 | 29.5 | Rep 1 | 143 | Rep 1 | 144
Rep 2 | 146 | Rep 2 | 146 | Rep 2 | 29.1 | Rep 2 | 29.5 | Rep 2 | 144 | Rep 2 | 143
Rep 3 | 146 | Rep 3 | 146 | Rep 3 | 29.4 | Rep 3 | 29.2 | Rep 3 | 143 | Rep 3 | 144
Rep 4 | 145 | Rep 4 | 146 | Rep 4 | 29.3 | Rep 4 | 29 | Rep 4 | 143 | Rep 4 | 142
Rep 5 | 146 | Rep 5 | 146 | Rep 5 | 29.9 | Rep 5 | 29.5 | Rep 5 | 144 | Rep 5 | 142
Target | 145.8 | % Diff. | 0.13717 | Target | 29.36 | % Diff. | 0.47684 | Target | 143.4 | % Diff. | 0.41841
(mean) |  | (Rep 1) |  | (mean) |  | (Rep 1) |  | (mean) |  | (Rep 1)
| Carrier | 146 |  |  | Carrier | 147 |  |  | Carrier | 148
Donor |  | Recipient |  | Donor |  | Recipient |  | Donor |  | Recipient
(Cup 4) | ACET | (Cup 5) | DIG | (Cup 4) | DIG | (Cup 5) | ACET | (Cup 4) | ACET | (Cup 5) | QUIN
Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results
Rep 1 | 2.9 | Rep 1 | 2.93 | Rep 1 | 144 | Rep 1 | 144 | Rep 1 | 4.9 | Rep 1 | 4.8
Rep 2 | 2.94 | Rep 2 | 2.92 | Rep 2 | 144 | Rep 2 | 144 | Rep 2 | 4.9 | Rep 2 | 4.8
Rep 3 | 2.93 | Rep 3 | 2.93 | Rep 3 | 144 | Rep 3 | 144 | Rep 3 | 4.8 | Rep 3 | 4.9
Rep 4 | 2.92 | Rep 4 | 2.91 | Rep 4 | 144 | Rep 4 | 143 | Rep 4 | 5 | Rep 4 | 4.8
Rep 5 | 2.88 | Rep 5 | 2.9 | Rep 5 | 143 | Rep 5 | 143 | Rep 5 | 4.8 | Rep 5 | 4.8
Target | 2.914 | % Diff. | 0.54907 | Target | 143.8 | % Diff. | 0.13908 | Target | 4.88 | % Diff. | -1.6393
(mean) |  | (Rep 1) |  | (mean) |  | (Rep 1) |  | (mean) |  | (Rep 1)
| Carrier | 149 |  |  | Carrier | 150 |  |  | Carrier | 151
Donor |  | Recipient |  | Donor |  | Recipient |  | Donor |  | Recipient
(Cup 4) | QUIN | (Cup 5) | ACET | (Cup 4) | ACET | (Cup 5) | CreaC | (Cup 4) | CreaC | (Cup 5) | ACET
Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results
Rep 1 | 144 | Rep 1 | 143 | Rep 1 | 7.19 | Rep 1 | 7.19 | Rep 1 | 113.1 | Rep 1 | 113
Rep 2 | 144 | Rep 2 | 144 | Rep 2 | 7.21 | Rep 2 | 7.16 | Rep 2 | 113.3 | Rep 2 | 112.9
Rep 3 | 144 | Rep 3 | 144 | Rep 3 | 7.11 | Rep 3 | 7.16 | Rep 3 | 113 | Rep 3 | 112.9
Rep 4 | 144 | Rep 4 | 143 | Rep 4 | 7.19 | Rep 4 | 7.16 | Rep 4 | 113 | Rep 4 | 112.1
Rep 5 | 143 | Rep 5 | 143 | Rep 5 | 7.27 | Rep 5 | 7.16 | Rep 5 | 112.4 | Rep 5 | 110.8
Target | 143.8 | % Diff. | -0.55633 | Target | 7.194 | % Diff. | -0.055602 | Target | 112.96 | % Diff. | 0.035411
(mean) |  | (Rep 1) |  | (mean) |  | (Rep 1) |  | (mean) |  | (Rep 1)

## CreaCU_PHNO_THEOP_PHNY_CARBA
| Carrier | 152 |  |  | Carrier | 153 |  |  | Carrier | 154
Donor |  | Recipient |  | Donor |  | Recipient |  | Donor |  | Recipient
(Cup 4) | ACET | (Cup 5) | CreaCU | (Cup 4) | CreaCU | (Cup 5) | ACET | (Cup 4) | ACET | (Cup 5) | PHNO
Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results
Rep 1 | 140.57 | Rep 1 | 140.43 | Rep 1 | 112.6 | Rep 1 | 112.9 | Rep 1 | 49.88 | Rep 1 | 50.89
Rep 2 | 140.54 | Rep 2 | 140.7 | Rep 2 | 113.1 | Rep 2 | 112.3 | Rep 2 | 50.16 | Rep 2 | 50.58
Rep 3 | 141.05 | Rep 3 | 138.67 | Rep 3 | 112.9 | Rep 3 | 112.4 | Rep 3 | 50.09 | Rep 3 | 50.12
Rep 4 | 138.76 | Rep 4 | 136.16 | Rep 4 | 112.8 | Rep 4 | 112 | Rep 4 | 49.62 | Rep 4 | 51.04
Rep 5 | 143.13 | Rep 5 | 138.42 | Rep 5 | 112.6 | Rep 5 | 112.4 | Rep 5 | 48.07 | Rep 5 | 50.77
Target | 140.81 | % Diff. | -0.26987 | Target | 112.8 | % Diff. | 0.088652 | Target | 49.564 | % Diff. | 2.6753
(mean) |  | (Rep 1) |  | (mean) |  | (Rep 1) |  | (mean) |  | (Rep 1)
| Carrier | 155 |  |  | Carrier | 156 |  |  | Carrier | 157
Donor |  | Recipient |  | Donor |  | Recipient |  | Donor |  | Recipient
(Cup 4) | PHNO | (Cup 5) | ACET | (Cup 4) | ACET | (Cup 5) | THEOP | (Cup 4) | THEOP | (Cup 5) | ACET
Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results
Rep 1 | 112.8 | Rep 1 | 112.8 | Rep 1 | 29.41 | Rep 1 | 29.45 | Rep 1 | 113.3 | Rep 1 | 112.7
Rep 2 | 112.9 | Rep 2 | 113 | Rep 2 | 29.18 | Rep 2 | 30.44 | Rep 2 | 113.1 | Rep 2 | 112.7
Rep 3 | 113 | Rep 3 | 112.8 | Rep 3 | 29.71 | Rep 3 | 29.52 | Rep 3 | 112.9 | Rep 3 | 112.8
Rep 4 | 113 | Rep 4 | 112.6 | Rep 4 | 30 | Rep 4 | 29.62 | Rep 4 | 112.6 | Rep 4 | 112.3
Rep 5 | 113.1 | Rep 5 | 112.8 | Rep 5 | 31.85 | Rep 5 | 30.08 | Rep 5 | 112.3 | Rep 5 | 112
Target | 112.96 | % Diff. | -0.14164 | Target | 30.03 | % Diff. | -1.9314 | Target | 112.84 | % Diff. | -0.12407
(mean) |  | (Rep 1) |  | (mean) |  | (Rep 1) |  | (mean) |  | (Rep 1)
| Carrier | 158 |  |  | Carrier | 159 |  |  | Carrier | 160
Donor |  | Recipient |  | Donor |  | Recipient |  | Donor |  | Recipient
(Cup 4) | ACET | (Cup 5) | PHNY | (Cup 4) | PHNY | (Cup 5) | ACET | (Cup 4) | ACET | (Cup 5) | CARBA
Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results
Rep 1 | 26.51 | Rep 1 | 26.39 | Rep 1 | 113.4 | Rep 1 | 113.9 | Rep 1 | 12.38 | Rep 1 | 12.41
Rep 2 | 26.12 | Rep 2 | 26.23 | Rep 2 | 113.3 | Rep 2 | 113.4 | Rep 2 | 12.41 | Rep 2 | 12.56
Rep 3 | 26.77 | Rep 3 | 26.67 | Rep 3 | 113.8 | Rep 3 | 113.1 | Rep 3 | 12.36 | Rep 3 | 12.39
Rep 4 | 26.77 | Rep 4 | 26.48 | Rep 4 | 113.3 | Rep 4 | 113.1 | Rep 4 | 12.37 | Rep 4 | 12.37
Rep 5 | 26.19 | Rep 5 | 26.46 | Rep 5 | 113.6 | Rep 5 | 112.7 | Rep 5 | 12.53 | Rep 5 | 12.42
Target | 26.472 | % Diff. | -0.30976 | Target | 113.48 | % Diff. | 0.37011 | Target | 12.41 | % Diff. | 0
(mean) |  | (Rep 1) |  | (mean) |  | (Rep 1) |  | (mean) |  | (Rep 1)

## CARBA_AlbP_Caffeine_PCPQ_CreaC
| Carrier | 161 |  |  | Carrier | 162 |  |  | Carrier | 163
Donor |  | Recipient |  | Donor |  | Recipient |  | Donor |  | Recipient
(Cup 4) | CARBA | (Cup 5) | ACET | (Cup 4) | ACET | (Cup 5) | AlbP | (Cup 4) | AlbP | (Cup 5) | ACET
Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results
Rep 1 | 113.6 | Rep 1 | 113.5 | Rep 1 | 3 | Rep 1 | 3 | Rep 1 | 115 | Rep 1 | 115
Rep 2 | 113.7 | Rep 2 | 113.4 | Rep 2 | 3 | Rep 2 | 3 | Rep 2 | 115 | Rep 2 | 115
Rep 3 | 113.9 | Rep 3 | 113.6 | Rep 3 | 3 | Rep 3 | 3 | Rep 3 | 115 | Rep 3 | 115
Rep 4 | 113.3 | Rep 4 | 112.9 | Rep 4 | 3 | Rep 4 | 3 | Rep 4 | 114 | Rep 4 | 114
Rep 5 | 113.6 | Rep 5 | 112.9 | Rep 5 | 3 | Rep 5 | 3 | Rep 5 | 113 | Rep 5 | 113
Target | 113.62 | % Diff. | -0.10562 | Target | 3 | % Diff. | 0 | Target | 114.4 | % Diff. | 0.52448
(mean) |  | (Rep 1) |  | (mean) |  | (Rep 1) |  | (mean) |  | (Rep 1)
| Carrier | 164 |  |  | Carrier | 165 |  |  | Carrier | 166
Donor |  | Recipient |  | Donor |  | Recipient |  | Donor |  | Recipient
(Cup 4) | ACET | (Cup 5) | Caffeine* | (Cup 4) | Caffeine* | (Cup 5) | ACET | (Cup 4) | ACET | (Cup 5) | PCPQ
Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results
Rep 1 | 15.5 | Rep 1 | 15.6 | Rep 1 | 115 | Rep 1 | 115 | Rep 1 | -1 | Rep 1 | -2
Rep 2 | 15.4 | Rep 2 | 15.3 | Rep 2 | 115 | Rep 2 | 115 | Rep 2 | -2 | Rep 2 | -2
Rep 3 | 14.9 | Rep 3 | 15.2 | Rep 3 | 114 | Rep 3 | 115 | Rep 3 | -2 | Rep 3 | -2
Rep 4 | 15.4 | Rep 4 | 15.2 | Rep 4 | 114 | Rep 4 | 114 | Rep 4 | -2 | Rep 4 | -2
Rep 5 | 15.4 | Rep 5 | 13.4 | Rep 5 | 113 | Rep 5 | 114 | Rep 5 | -2 | Rep 5 | -1
Target | 15.32 | % Diff. | 1.8277 | Target | 114.2 | % Diff. | 0.70053 | Target | 0 | % Diff. | N/A
(mean) |  | (Rep 1) |  | (mean) |  | (Rep 1) |  | (mean) |  | (Rep 1)
| Carrier | 167 |  |  | Carrier | 168 |  |  | Carrier | 169
Donor |  | Recipient |  | Donor |  | Recipient |  | Donor |  | Recipient
(Cup 4) | PCPQ | (Cup 5) | ACET | (Cup 4) | ACET | (Cup 5) | CreaC | (Cup 4) | CreaC | (Cup 5) | ACET
Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results
Rep 1 | 115 | Rep 1 | 115 | Rep 1 | 1 | Rep 1 | 1 | Rep 1 | 115 | Rep 1 | 115
Rep 2 | 115 | Rep 2 | 115 | Rep 2 | 1.1 | Rep 2 | 1 | Rep 2 | 115 | Rep 2 | 114
Rep 3 | 114 | Rep 3 | 115 | Rep 3 | 1 | Rep 3 | 1 | Rep 3 | 115 | Rep 3 | 115
Rep 4 | 114 | Rep 4 | 114 | Rep 4 | 1 | Rep 4 | 1 | Rep 4 | 114 | Rep 4 | 113
Rep 5 | 113 | Rep 5 | 114 | Rep 5 | 1 | Rep 5 | 1 | Rep 5 | 113 | Rep 5 | 112
Target | 114.2 | % Diff. | 0.70053 | Target | 1.02 | % Diff. | -1.9608 | Target | 114.4 | % Diff. | 0.52448
(mean) |  | (Rep 1) |  | (mean) |  | (Rep 1) |  | (mean) |  | (Rep 1)
|  |  | *Non-Abbott application |  |  |  |  | N/A due to negative readings for Cup 2 & 5

## A1AT_B2M_IgE
| Carrier | 170 |  |  | Carrier | 171 |  |  | Carrier | 172
Donor |  | Sheet 20 |  | Donor |  | Recipient |  | Donor |  | Recipient
(Cup 4) | ACET | (Cup 5) | A1AT | (Cup 4) | A1AT | (Cup 5) | ACET | (Cup 4) | ACET | (Cup 5) | B2M
Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results
Rep 1 | 168 | Rep 1 | 166 | Rep 1 | 116 | Rep 1 | 115 | Rep 1 | 3.67 | Rep 1 | 3.66
Rep 2 | 166 | Rep 2 | 166 | Rep 2 | 116 | Rep 2 | 116 | Rep 2 | 3.67 | Rep 2 | 3.65
Rep 3 | 166 | Rep 3 | 167 | Rep 3 | 116 | Rep 3 | 115 | Rep 3 | 3.66 | Rep 3 | 3.67
Rep 4 | 166 | Rep 4 | 167 | Rep 4 | 115 | Rep 4 | 115 | Rep 4 | 3.66 | Rep 4 | 3.65
Rep 5 | 167 | Rep 5 | 166 | Rep 5 | 116 | Rep 5 | 115 | Rep 5 | 3.75 | Rep 5 | 3.68
Target | 166.6 | % Diff. | -0.36014 | Target | 115.8 | % Diff. | -0.69085 | Target | 3.682 | % Diff. | -0.5975
(mean) |  | (Rep 1) |  | (mean) |  | (Rep 1) |  | (mean) |  | (Rep 1)
| Carrier | 173 |  |  | Carrier | 174 |  |  | Carrier | 175
Donor |  | Recipient |  | Donor |  | Recipient |  | Donor |  | Recipient
(Cup 4) | B2M | (Cup 5) | ACET | (Cup 4) | ACET | (Cup 5) | IgE | (Cup 4) | IgE | (Cup 5) | ACET
Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results
Rep 1 | 116 | Rep 1 | 116 | Rep 1 | 166 | Rep 1 | 166.5 | Rep 1 | 116 | Rep 1 | 116
Rep 2 | 116 | Rep 2 | 115 | Rep 2 | 181.1 | Rep 2 | 167.6 | Rep 2 | 116 | Rep 2 | 116
Rep 3 | 115 | Rep 3 | 116 | Rep 3 | 174.2 | Rep 3 | 174.4 | Rep 3 | 116 | Rep 3 | 116
Rep 4 | 114 | Rep 4 | 115 | Rep 4 | 171.3 | Rep 4 | 177.6 | Rep 4 | 115 | Rep 4 | 115
Rep 5 | 114 | Rep 5 | 116 | Rep 5 | 166 | Rep 5 | 170.4 | Rep 5 | 116 | Rep 5 | 115
Target | 115 | % Diff. | 0.86957 | Target | 171.72 | % Diff. | -3.0398 | Target | 115.8 | % Diff. | 0.17271
(mean) |  | (Rep 1) |  | (mean) |  | (Rep 1) |  | (mean) |  | (Rep 1)

## PrpxQ_Ethylene_Bupre_Oxycod
| Carrier | 179 |  |  | Carrier | 180 |  |  | Carrier | 181
Donor |  | Shee |  | Donor |  | Recipient |  | Donor |  | Recipient
(Cup 4) | ACET | (Cup 5) | PrpxQ | (Cup 4) | PrpxQ | (Cup 5) | ACET | (Cup 4) | ACET | (Cup 5) | Ethlyene Glycol* | l
Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results
Rep 1 | 316 | Rep 1 | 310 | Rep 1 | 114 | Rep 1 | 113 | Rep 1 | 242 | Rep 1 | 235
Rep 2 | 315 | Rep 2 | 313 | Rep 2 | 113 | Rep 2 | 113 | Rep 2 | 241 | Rep 2 | 238
Rep 3 | 317 | Rep 3 | 310 | Rep 3 | 114 | Rep 3 | 113 | Rep 3 | 240 | Rep 3 | 236
Rep 4 | 316 | Rep 4 | 306 | Rep 4 | 114 | Rep 4 | 114 | Rep 4 | 240 | Rep 4 | 239
Rep 5 | 317 | Rep 5 | 309 | Rep 5 | 114 | Rep 5 | 113 | Rep 5 | 243 | Rep 5 | 241
Target | 316.2 | % Diff. | -1.9608 | Target | 113.8 | % Diff. | -0.70299 | Target | 241.2 | % Diff. | -2.5705
(mean) |  | (Rep 1) |  | (mean) |  | (Rep 1) |  | (mean) |  | (Rep 1)
| Carrier | 182 |  |  | Carrier | 183 |  |  | Carrier | 184
Donor |  | Recipient |  | Donor |  | Recipient |  | Donor |  | Recipient
(Cup 4) | Ethlyene Glycol* | (Cup 5) | ACET | (Cup 4) | ACET | (Cup 5) | Buprenorph* | (Cup 4) | Buprenorph* | (Cup 5) | ACET
Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results
Rep 1 | 141 | Rep 1 | 141 | Rep 1 | 7 | Rep 1 | 8 | Rep 1 | 141 | Rep 1 | 143
Rep 2 | 142 | Rep 2 | 141 | Rep 2 | 8 | Rep 2 | 8 | Rep 2 | 143 | Rep 2 | 143
Rep 3 | 143 | Rep 3 | 141 | Rep 3 | 8 | Rep 3 | 9 | Rep 3 | 142 | Rep 3 | 143
Rep 4 | 142 | Rep 4 | 143 | Rep 4 | 8 | Rep 4 | 8 | Rep 4 | 143 | Rep 4 | 143
Rep 5 | 141 | Rep 5 | 142 | Rep 5 | 8 | Rep 5 | 9 | Rep 5 | 144 | Rep 5 | 142
Target | 141.8 | % Diff. | -0.56417 | Target | 7.8 | % Diff. | 2.5641 | Target | 142.6 | % Diff. | 0.2805
(mean) |  | (Rep 1) |  | (mean) |  | (Rep 1) |  | (mean) |  | (Rep 1)
| Carrier | 185 |  |  | Carrier | 186
Donor |  | Recipient |  | Donor |  | Recipient
(Cup 4) | ACET | (Cup 5) | Oxycodone* | (Cup 4) | Oxycodone* | (Cup 5) | ACET
Cup 2 Results |  | Cup 5 Results |  | Cup 2 Results |  | Cup 5 Results
Rep 1 | 122 | Rep 1 | 123 | Rep 1 | 143 | Rep 1 | 142
Rep 2 | 121 | Rep 2 | 124 | Rep 2 | 142 | Rep 2 | 141
Rep 3 | 124 | Rep 3 | 123 | Rep 3 | 143 | Rep 3 | 144
Rep 4 | 122 | Rep 4 | 124 | Rep 4 | 143 | Rep 4 | 142
Rep 5 | 125 | Rep 5 | 120 | Rep 5 | 144 | Rep 5 | 143
Target | 122.8 | % Diff. | 0.16287 | Target | 143 | % Diff. | -0.6993
(mean) |  | (Rep 1) |  | (mean) |  | (Rep 1)
*Non-Abbott application